## Slide 1: Super Fast Automated Emulsion Scanning System ~ HTS ~
Masahiro Yoshimoto,
Toshiyuki Nakano and other HTS collaboration
Nagoya University, Japan

## Slide 2: Our Emulsion Film
Before development
After development
New Emulsion film (After development)
OPERA film
In Nagoya Univ., We have been developing new emulsion film.
Grain density is much higher.
Emulsion gel can be produced by ourselves.
Pouring (coating) facility is under construction.
12.5cm×10.0cm

## Slide 3: Detecting ionizing particles
Image sensor
Objective lens
Emulsion layer
Plastic base
Emulsion layer
Condenser lens
Light source
Automated stage
Emulsion plate
Kohler illumination
Charged particles

## Slide 4: Capturing Emulsion Image
Thickness of the emulsion layer >50mm
Depth of field ~3mm
Size of the silver grain ~0.3mm
Optical resolution which is required <0.5mm
1μｍ
1μｍ
Stacked image of optical microscope
res. 0.27μm
Image of X-ray microscope
DOF is ±~70μm
res. 0.07 μｍ
in focus
out focus
out focus
α-ray track
Z
OPERA film

## Slide 5: Image processing
Raw image which is taken

## Slide 6: Image processing
Applying 2D-FIR filter
Cut off the background

## Slide 7
Track recognition method
Take 16 tomographic images by microscope optics.
Shift images to aim at specific angle tracks.
Sum up 16 images to examine coincidence.
Find signal of tracks.
Repeat for all angles in space, >2000 times
Invented by K. Niwa in 1974

## Slide 8: History of emulsion readout system in Nagoya
TS (TTL)
1983~
CHORUS(R&D)
0.003cm2/h
NTS (CPLD) 1994~
CHORUS
0.082cm2/h
UTS (FPGA) 1998~
DONUT
1cm2/h
Innovation has been required for the next generation systems.

## Slide 9: Evolution of the Scanning Speed
10000
1000
100
10
1
0.1
0.01
TS(TTL)
1983
NTS(CPLD)
1994
UTS(FPGA)
1998
S-UTS
(FPGA)
2006-
Running
0.003
0.082
1
72
0.001
Speed in cm2/ hour
×100

## Slide 10: Overcome the Bottle necks of the image acquisition
Use Ultra High Speed Camera
  Up to 3k frames per second.
  60 view/sec (20times)
Image taking by follow shot
  No step and repeat operation can avoid a mechanical bottleneck.
  High speed motion of the objective lens can be done by piezo actuators
Optimizing Field of View
  120mm×90mm -> 200mm×200mm (4 times)
We have achieved 72cm2/h

## Slide 11: Movie of the readout by S-UTS
S-UTS

## Slide 12: S-UTS (Super-Ultra Track Selector)
Five S-UTSs were produced in order to perform OPERA experiment.
Readout capability of one S-UTS is 20m2/year.

## Slide 13: Applications of Nuclear Emulsion ・・・・・ need much more speed !
Muon Radiography
cosmic ray muon
100m2 will be required for one measurement.
→Morishima’s talk

## Slide 14: Applications of Nuclear Emulsion ・・・・・ need much more speed !
Several 1000m2 will be required for scientific observation.
GRAINE project
Gamma-ray Telescope
2011 Balloon experiment was performed.
2014 Planned at Alice Springs.
2015- Starting scientific observation
Timestamper
Attitude monitor
e+
e-
g
Calorimeter
Converter

## Slide 15: Applications of Nuclear Emulsion ・・・・・ need much more speed !
Dark Matter Search
Dark Matter
Sun
230km/sec
F.L.～ 200nm
NIT
WIMP
Recoil atoms
gelatin H,C,N,O,S
Earth
Vrecoil = 2 ( Vsun + VWIMP )
= 100～1000km/sec
crystal Ag,Br,I
We need to readout the target mass of 1 ton (about 4000m2).
→Asada’s talk
~3g / film
0.0125m2/film

## Slide 16: Evolution of the Scanning Speed
※1 Area of each layer
※2 Area of the films
10000
1000
100
10
1
0.1
0.01
TS(TTL)
1983
NTS(CPLD)
1994
UTS(FPGA)
1998
S-UTS
(FPGA)
2006-
Running
HTS
(GPGPU)
2011-
Developing
0.003
0.082
1
72
9000
0.001
Speed in cm2/ hour
×100
~2000m2/year
~20m2/year
※2
※1

## Slide 17: Concept of HTS (Hyper Track Selector)
It is difficult to drive objective much more quickly.
Field of view can be a few orders of magnitude larger.

## Slide 18
Computer
GPU board
・Wide FOV optical system
・High precision and speed stage
X
Y
Z
HTS

## Slide 19: Enormous Objective
Resolution : ~420nm
N.A. : 0.65
Light source : g-line (436nm) Magnitude : ×12.2
F.O.V : 5.1 (H)×5.1 (V)mm
#of image plane 6
(by Beam splitter)
Weight : 80kg

## Slide 20: Wide field of view
Field of view (FOV) of HTS
5mm
~0.2mm
5mm
× 600
S-UTS

## Slide 21: Speed and Coverage of Mosaic Imager
Divide FOV into 72 parts.
Need the sensor of 2M pixel and 340fps.
Specially ordered Mosaic Imager
5mm
5mm

## Slide 22: Mosaic Imager system
1
2
3
4
5
6
Objective lens
Illumination
Beam splitter
1
Camera1
Camera2
Camera6
5mm
5mm

## Slide 23: Movie of the readout by HTS
HTS

## Slide 24: Image capturing at rest
Stopping accuracy should be 0.3μm.
HTS takes advantage of Stop & Go in image capturing.
Key is how to stop quickly and precisely.
The Z-axis should move at constant velocity.
Throughput = FOV × Repetition freq.
FOV : 600 times larger
To achieve 100 times faster than S-UTS,
Repetition freq. should be >10view/s (1/6 of SUTS)

## Slide 25: High precision and speed stage
Reaction force is canceled by counter stage
Counter Stage
Main Stage
capture (18frame)
20ms
20ms
20ms
drive
Time
Velocity
Moving mass : 34.7kg
Step dist. : 5.0mm
Acc. rate : 5G (50m/s2)
Acpt. vibration.: ±0.3mm
50cm/sec (Max. speed)
100ms (10view/s)
60ms
settle
Under developing

## Slide 26: Evaluation of stage settling time
By optimizing the acceleration, 5view/sec was achieved.
The displacement has been measured by using main optics.
Chart: 
displacement [micron]
±0.25μm
msec
capture
60ms
0.50
0.25
-0.50
-0.75
0
-0.25
-1.00
-1.25
drive 120ms
settle
17ms

## Slide 27: Tracking Hardware
Sensor 72 × 36Mbyte/view × 10view/s ＝ 26GB/s
Sensor
Computer
GPU
Frame Grabber
Frame Grabber
CPU
GPU
×36
360MB/s
360MB/s
26GB/s/all
2.6TB/plate (OPERA film)
Computer 36, Frame grabber board 72, GPU board72

## Slide 28: Dataflow
Sensor
Image filter
Track recognition
Clustering
Storage Server
36MB
GPU 15~50ms
GPU 20~50ms
depend on the angle range.
CPU 100~200ms
Raw Image
All track
Micro Track
~0.3MB
Binarized Image
5MB
Amount of data
per sensor per view
5MB
Processing time
Synchronous
Asynchronous
Total is less than 100ms
pixel count

## Slide 29: Stage control
2Sensor & 1Computer
Image Capture &Tracking
Stage controller
・Motor (8 axis)
・Piezo (3 actuator)
Master Computer
Image Capture &Tracking
10G Ethernet Switch
Storage Server
10G Ether
RAID DISKs
Broadcast stage position, actuator status and Trigger timing for image capture
Feedback Pixel count for surface detection
6MB/s
・・・
ave. 200MB/s/all
~10GB/plate
×36
Storage Server

## Slide 30: Scanning speed and other status
December 2011 Objective lens and beam splitter were mounted.
May 2012 All motors for xyz-axis drive began to work.
November 2012 Piezo actuator for z-axis drive was installed.
May 2013 Half of cameras and computers became ready.
May 2013~ Scanning test for various emulsion plates
 | Practicable | Installed | Final Goal
Stage | 2 view/s | 5 view/s | 10 view/s
Camera | 1 unit | 3 unit | 6 unit
Speed | 300cm2/h | 2250cm2/h | 9000cm2/h
cf. S-UTS 72cm2/h
I did scanning test at this speed.

## Slide 31: New Emulsion Film
Efficiency 98.0%
18,264track / 18,632track
tanθ < 0.6
10
20
30
[deg]
HTS
PL13
PL11
Exist or not
？
cf.
S-UTS 98.3%
New type
GD=51.6+-2.6
FD=3.3+-0.4
preliminary

## Slide 32: New Emulsion Film
Angle deviation
PL13
PL11
prediction
δθ
Sigma
4.5mrad
Sigma
5.2mrad
X-projection
Y-projection
cf. S-UTS
4.8mrad
cf. S-UTS
4.3mrad
preliminary

## Slide 33
What we can do with HTS?

## Slide 34: Emulsion Archives Project
The project is to read all emulsion and convert to the data.
In near future we will scan a lot of emulsion with HTS.
Requirement,
All Emulsion existing around the world.
Money (and man power) for running
Unified Specification of track and image
And
Dream for Science

## Slide 35: Summary and Prospect
We are developing HTS which is 100 times faster.
Scanning at the speed of 300cm2/h is practicable.
The efficiency and angle resolution are almost the same as previous system.
HTS is working for muon radiography, GRAINE project and dark matter search.
All cameras and computers will be installed this year.